## Sheet1
Date | Cal # | Fluorometer # | Work Area | Point # | Solution | Chl Std. Conc. | ml of standard in 10ml acetone | volume of acetone (ml) | Total Volume | chl standard concentration | Blank Before Acid | Blank After Acid | Raw Before Acid | Rb (Blank Subtracted) | Raw After Acid | Ra (Blank Subtracted) | FM (Blank Subtracted) | Calc. Chl | Calc. Phaeo | Avg (standard and estimated) | Stdev. | COV | lab | Comments
2015-03-19 | Q1 | 720001154 | Quadra | 1 | ND | ND | ND | ND | ND | 0.203 | 1314.05 |  | 2131.4 | 817.35 | 1903.06 | 589.01 | 1.3877 | 0.24071 | 0.21046 | 0.22185 | 0.018853 | 0.084978 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 2 | ND | ND | ND | ND | ND | 0.5075 | 1314.05 |  | 2201.59 | 887.54 | 1874.99 | 560.94 | 1.5822 | 0.34429 | 0.085378 | 0.42589 | 0.081607 | 0.19161 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 3 | ND | ND | ND | ND | ND | 1.015 | 1314.05 |  | 3452.54 | 2138.49 | 2538.29 | 1224.24 | 1.7468 | 0.96376 | -0.026026 | 0.98938 | 0.02562 | 0.025895 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 4 | ND | ND | ND | ND | ND | 2.03 | 1314.05 |  | 4411.22 | 3097.17 | 2911.07 | 1597.02 | 1.9393 | 1.5814 | -0.35812 | 1.8057 | 0.22431 | 0.12422 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 5 | ND | ND | ND | ND | ND | 4.06 | 1314.05 |  | 9224.24 | 7910.19 | 6045.21 | 4731.16 | 1.6719 | 3.3512 | 0.27275 | 3.7056 | 0.35441 | 0.095641 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 6 | ND | ND | ND | ND | ND | 6.09 | 1314.05 |  | 13798.43 | 12484.38 | 8608.41 | 7294.36 | 1.7115 | 5.4711 | 0.1162 | 5.7805 | 0.30946 | 0.053535 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 7 | ND | ND | ND | ND | ND | 8.12 | 1314.05 |  | 17940.5 | 16626.45 | 11146.67 | 9832.62 | 1.6909 | 7.1617 | 0.36977 | 7.6409 | 0.47913 | 0.062706 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 8 | ND | ND | ND | ND | ND | 12.18 | 1314.05 |  | 26258.75 | 24944.7 | 15984.26 | 14670.21 | 1.7004 | 10.8309 | 0.40608 | 11.5054 | 0.67456 | 0.058629 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 9 | ND | ND | ND | ND | ND | 18.44 | 1314.05 |  | 40513.64 | 39199.59 | 22800 | 21485.95 | 1.8244 | 18.6729 | -2.2153 | 18.5564 | 0.11645 | 0.0062752 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 10 | ND | ND | ND | ND | ND | 36.88 | 1314.05 |  | 82378.18 | 81064.13 | 47841.26 | 46527.21 | 1.7423 | 36.4072 | -0.76868 | 36.6436 | 0.2364 | 0.0064513 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 11 | ND | ND | ND | ND | ND | 55.32 | 1314.05 |  | 122535.33 | 121221.28 | 71708.54 | 70394.49 | 1.722 | 53.5792 | 0.34097 | 54.4496 | 0.87039 | 0.015985 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 12 | ND | ND | ND | ND | ND | 73.76 | 1314.05 |  | 165425.87 | 164111.82 | 96826.49 | 95512.44 | 1.7182 | 72.3143 | 0.84557 | 73.0371 | 0.72287 | 0.0098973 | Hakai
2015-03-19 | Q1 | 720001154 | Quadra | 13 | ND | ND | ND | ND | ND | 92.2 | 1314.05 |  | 209282.28 | 207968.23 | 120994 | 119679.95 | 1.7377 | 93.0694 | -1.3979 | 92.6347 | 0.43469 | 0.0046925 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 1 | Solution C | 22.8 | 0.1 | 9.9 | 10 | 0.228 | 1916.88 |  | 2161.96 | 245.08 | 1932.96 | 16.08 | 15.2413 | 0.28996 | -0.27634 | 0.25898 | 0.03098 | 0.11962 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 2 | Solution C | 22.8 | 0.25 | 9.75 | 10 | 0.57 | 1916.88 |  | 2556.76 | 639.88 | 1962.62 | 45.74 | 13.9895 | 0.7523 | -0.71356 | 0.66115 | 0.091149 | 0.13787 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 3 | Solution C | 22.8 | 0.5 | 9.5 | 10 | 1.14 | 1916.88 |  | 3600.28 | 1683.4 | 2695.44 | 778.56 | 2.1622 | 1.1457 | -0.48631 | 1.1429 | 0.0028533 | 0.0024966 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 4 | Solution C | 22.8 | 1 | 9 | 10 | 2.28 | 1916.88 |  | 5454.74 | 3537.86 | 3789.78 | 1872.9 | 1.889 | 2.1082 | -0.52192 | 2.1941 | 0.085916 | 0.039158 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 5 | Solution C | 22.8 | 2 | 8 | 10 | 4.56 | 1916.88 |  | 5075.08 |  | 3428.95 | 1512.07 | 2.0887 |  | -0.80368 | 4.56 |  |  | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 6 | Solution C | 22.8 | 3 | 7 | 10 | 6.84 | 1916.88 |  | 14361.49 | 12444.61 | 9431.34 | 7514.46 | 1.6561 | 6.2425 | 0.1218 | 6.5413 | 0.29873 | 0.045668 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 7 | Solution C | 22.8 | 4 | 6 | 10 | 9.12 | 1916.88 |  | 18527.41 | 16610.53 | 11659.47 | 9742.59 | 1.7049 | 8.6962 | -0.44472 | 8.9081 | 0.21191 | 0.023789 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 8 | Solution C | 22.8 | 6 | 4 | 10 | 13.68 | 1916.88 |  | 27129.61 | 25212.73 | 16904.44 | 14987.56 | 1.6822 | 12.9471 | -0.25342 | 13.3135 | 0.36646 | 0.027525 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 9 | Solution B | 216.8 | 1 | 9 | 10 | 21.68 | 1916.88 |  | 40237.17 | 38320.29 | 25257.03 | 23340.15 | 1.6418 | 18.9678 | 0.80004 | 20.3239 | 1.3561 | 0.066724 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 10 | Solution B | 216.8 | 2 | 8 | 10 | 43.36 | 1916.88 |  | 85612.06 | 83695.18 | 51937.76 | 50020.88 | 1.6732 | 42.6383 | -0.27331 | 42.9992 | 0.36084 | 0.0083917 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 11 | Solution B | 216.8 | 3 | 7 | 10 | 65.04 | 1916.88 |  | 128726.41 | 126809.53 | 77723.59 | 75806.71 | 1.6728 | 64.5797 | -0.37542 | 64.8098 | 0.23017 | 0.0035514 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 12 | Solution B | 216.8 | 4 | 6 | 10 | 86.72 | 1916.88 |  | 170417.14 | 168500.26 | 102577.59 | 100660.71 | 1.6739 | 85.8983 | -0.64404 | 86.3091 | 0.41085 | 0.0047602 | Hakai
2015-08-06 | Q2 | 720001154 | Quadra | 13 | Solution B | 216.8 | 5 | 5 | 10 | 108.4 | 1916.88 |  | 214818.29 | 212901.41 | 131255.46 | 129338.58 | 1.6461 | 105.8071 | 3.7358 | 107.1035 | 1.2965 | 0.012105 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 1 | Solution C | 18.3 | 0.1 | 9.9 | 10 | 0.183 | 1487.93 | 1529.27 | 1985.9 | 497.97 | 1713.07 | 183.8 | 2.7093 | 0.36952 | -0.22665 | 0.27626 | 0.093261 | 0.33758 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 2 | Solution C | 18.3 | 0.25 | 9.75 | 10 | 0.4575 | 1487.93 | 1529.27 | 2718.44 | 1230.51 | 2341.85 | 812.58 | 1.5143 | 0.49156 | 0.14008 | 0.47453 | 0.017031 | 0.035891 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 3 | Solution C | 18.3 | 0.5 | 9.5 | 10 | 0.915 | 1487.93 | 1529.27 | 4010.67 | 2522.74 | 3125.08 | 1595.81 | 1.5809 | 1.0902 | 0.15022 | 1.0026 | 0.08762 | 0.087391 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 4 | Solution C | 18.3 | 1 | 9 | 10 | 1.83 | 1487.93 | 1529.27 | 6004.74 | 4516.81 | 4597.64 | 3068.37 | 1.4721 | 1.7036 | 0.68149 | 1.7668 | 0.063184 | 0.035761 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 5 | Solution C | 18.3 | 2 | 8 | 10 | 3.66 | 1487.93 | 1529.27 | 10288.27 | 8800.34 | 6838.81 | 5309.54 | 1.6575 | 4.1058 | 0.021411 | 3.8829 | 0.22291 | 0.057408 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 6 | Solution C | 18.3 | 3 | 7 | 10 | 5.49 | 1487.93 | 1529.27 | 14331.59 | 12843.66 | 9363.97 | 7834.7 | 1.6393 | 5.8915 | 0.19864 | 5.6907 | 0.20073 | 0.035273 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 7 | Solution C | 18.3 | 4 | 6 | 10 | 7.32 | 1487.93 | 1529.27 | 17654.59 | 16166.66 | 11650.41 | 10121.14 | 1.5973 | 7.1106 | 0.75676 | 7.2153 | 0.10468 | 0.014508 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 8 | Solution C | 18.3 | 6 | 4 | 10 | 10.98 | 1487.93 | 1529.27 | 27090.43 | 25602.5 | 17164.73 | 15635.46 | 1.6375 | 11.7231 | 0.43075 | 11.3515 | 0.37154 | 0.03273 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 9 | Solution B | 201 | 1 | 9 | 10 | 20.1 | 1487.93 | 1529.27 | 44514.52 | 43026.59 | 26862.48 | 25333.21 | 1.6984 | 20.8107 | -1.1186 | 20.4553 | 0.35534 | 0.017371 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 10 | Solution B | 201 | 2 | 8 | 10 | 40.2 | 1487.93 | 1529.27 | 86622.39 | 85134.46 | 52929.49 | 51400.22 | 1.6563 | 39.6777 | 0.27697 | 39.9388 | 0.26116 | 0.006539 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 11 | Solution B | 201 | 3 | 7 | 10 | 60.3 | 1487.93 | 1529.27 | 129279.47 | 127791.54 | 77697.24 | 76167.97 | 1.6778 | 60.7188 | -1.5116 | 60.5094 | 0.20941 | 0.0034608 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 12 | Solution B | 201 | 4 | 6 | 10 | 80.4 | 1487.93 | 1529.27 | 175624.95 | 174137.02 | 107109.88 | 105580.61 | 1.6493 | 80.635 | 1.4354 | 80.5175 | 0.11748 | 0.0014591 | Hakai
2016-05-04 | Q3 | 720001154 | Quadra | 13 | Solution B | 201 | 5 | 5 | 10 | 100.5 | 1487.93 | 1529.27 | 217038.78 | 215550.85 | 130522.67 | 128993.4 | 1.671 | 101.8075 | -1.5378 | 101.1538 | 0.65375 | 0.0064629 | Hakai
2017-01-08 | Q4 | 720001154 | Quadra | 1 | Solution C | 18.3 | 0.1 | 9.9 | 10 | 0.183 | 1881.76 | 1748.57 | 2091.14 | 209.38 | 1933.3 | 184.73 | 1.1334 | 0.027683 | -2.0048 | 0.10534 | 0.077658 | 0.73721 | Hakai | FM (no blank) LOW
2017-01-08 | Q4 | 720001154 | Quadra | 2 | Solution C | 18.3 | 0.25 | 9.75 | 10 | 0.4575 | 1881.76 | 1748.57 | 2869.73 | 987.97 | 2532.9 | 784.33 | 1.2596 | 0.2287 | -1.7639 | 0.3431 | 0.1144 | 0.33344 | Hakai | FM (no blank) LOW
2017-01-08 | Q4 | 720001154 | Quadra | 3 | Solution C | 18.3 | 0.5 | 9.5 | 10 | 0.915 | 1881.76 | 1748.57 | 3804.79 | 1923.03 | 2911.79 | 1163.22 | 1.6532 | 0.8533 | -2.1092 | 0.88415 | 0.03085 | 0.034892 | Hakai | FM (no blank) LOW
2017-01-08 | Q4 | 720001154 | Quadra | 4 | Solution C | 18.3 | 1 | 9 | 10 | 1.83 | 1881.76 | 1748.57 | 5881.29 | 3999.53 | 3968.8 | 2220.23 | 1.8014 | 1.9982 | -2.475 | 1.9141 | 0.084116 | 0.043945 | Hakai | FM (no blank) LOW
2017-01-08 | Q4 | 720001154 | Quadra | 5 | Solution C | 18.3 | 2 | 8 | 10 | 3.66 | 1881.76 | 1748.57 | 10164.39 | 8282.63 | 6839.32 | 5090.75 | 1.627 | 3.5846 | -1.9456 | 3.6223 | 0.037689 | 0.010405 | Hakai | FM (no blank) LOW
2017-01-08 | Q4 | 720001154 | Quadra | 6 | Solution C | 18.3 | 3 | 7 | 10 | 5.49 | 1881.76 | 1748.57 | 14380.34 | 12498.58 | 9332.58 | 7584.01 | 1.648 | 5.5193 | -2.0425 | 5.5046 | 0.01464 | 0.0026596 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 7 | Solution C | 18.3 | 4 | 6 | 10 | 7.32 | 1881.76 | 1748.57 | 18445.4 | 16563.64 | 12033.53 | 10284.96 | 1.6105 | 7.0512 | -1.5836 | 7.1856 | 0.13438 | 0.018702 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 8 | Solution C | 18.3 | 6 | 4 | 10 | 10.98 | 1881.76 | 1748.57 | 27173.76 | 25292 | 16903.56 | 15154.99 | 1.6689 | 11.3843 | -2.327 | 11.1822 | 0.20216 | 0.018078 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 9 | Solution B | 201 | 1 | 9 | 10 | 20.1 | 1881.76 | 1748.57 | 44264.51 | 42382.75 | 27252.54 | 25503.97 | 1.6618 | 18.9556 | -2.2702 | 19.5278 | 0.57219 | 0.029301 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 10 | Solution B | 201 | 2 | 8 | 10 | 40.2 | 1881.76 | 1748.57 | 88554.36 | 86672.6 | 54005.09 | 52256.52 | 1.6586 | 38.6508 | -2.2464 | 39.4254 | 0.77461 | 0.019647 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 11 | Solution B | 201 | 3 | 7 | 10 | 60.3 | 1881.76 | 1748.57 | 131592.87 | 129711.11 | 79332.03 | 77583.46 | 1.6719 | 58.5417 | -3.469 | 59.4208 | 0.87917 | 0.014796 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 12 | Solution B | 201 | 4 | 6 | 10 | 80.4 | 1881.76 | 1748.57 | 179339.09 | 177457.33 | 113049.25 | 111300.68 | 1.5944 | 74.2968 | 5.6284 | 77.3484 | 3.0516 | 0.039452 | Hakai | FM (no blank) - Inconsistent Values
2017-01-08 | Q4 | 720001154 | Quadra | 13 | Solution B | 201 | 5 | 5 | 10 | 100.5 | 1881.76 | 1748.57 | 232313.32 | 230431.56 | 136980.17 | 135231.6 | 1.704 | 106.9138 | -9.3493 | 103.7069 | 3.2069 | 0.030923 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 1 | Solution C | 20 | 0.1 | 9.9 | 10 | 0.2 | 1600.08 | 1777.6 | 2180.59 | 580.51 | 2041.37 | 263.77 | 2.2008 | 0.42438 | -0.18564 | 0.31219 | 0.11219 | 0.35936 | Hakai | FM (no blank) LOW
2017-08-16 | Q5 | 720001154 | Quadra | 2 | Solution C | 20 | 0.25 | 9.75 | 10 | 0.5 | 1600.08 | 1777.6 | 2531.65 | 931.57 | 2307.58 | 529.98 | 1.7577 | 0.53806 | -0.058374 | 0.51903 | 0.01903 | 0.036665 | Hakai | FM (no blank) LOW
2017-08-16 | Q5 | 720001154 | Quadra | 3 | Solution C | 20 | 0.5 | 9.5 | 10 | 1 | 1600.08 | 1777.6 | 3518.03 | 1917.95 | 2905.04 | 1127.44 | 1.7012 | 1.0591 | -0.038695 | 1.0296 | 0.029573 | 0.028724 | Hakai | FM (no blank) LOW
2017-08-16 | Q5 | 720001154 | Quadra | 4 | Solution C | 20 | 1 | 9 | 10 | 2 | 1600.08 | 1777.6 | 4527.2 | 2927.12 | 3313.45 | 1535.85 | 1.9059 | 1.8641 | -0.47396 | 1.932 | 0.06797 | 0.03518 | Hakai | FM (no blank) LOW
2017-08-16 | Q5 | 720001154 | Quadra | 5 | Solution C | 20 | 2 | 8 | 10 | 4 | 1600.08 | 1777.6 | 8722.76 | 7122.68 | 6060.37 | 4282.77 | 1.6631 | 3.805 | 0.071367 | 3.9025 | 0.097506 | 0.024986 | Hakai | FM (no blank) LOW
2017-08-16 | Q5 | 720001154 | Quadra | 6 | Solution C | 20 | 3 | 7 | 10 | 6 | 1600.08 | 1777.6 | 13273.93 | 11673.85 | 8625.85 | 6848.25 | 1.7046 | 6.4655 | -0.26709 | 6.2327 | 0.23273 | 0.037341 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 7 | Solution C | 20 | 4 | 6 | 10 | 8 | 1600.08 | 1777.6 | 16652.4 | 15052.32 | 10802.13 | 9024.53 | 1.6679 | 8.0762 | 0.091948 | 8.0381 | 0.038097 | 0.0047395 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 8 | Solution C | 20 | 6 | 4 | 10 | 12 | 1600.08 | 1777.6 | 24612.37 | 23012.29 | 15530.55 | 13752.95 | 1.6733 | 12.4059 | 0.041945 | 12.203 | 0.20296 | 0.016632 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 9 | Solution B | 224 | 1 | 9 | 10 | 22.4 | 1600.08 | 1777.6 | 38964.1 | 37364.02 | 24478.22 | 22700.62 | 1.6459 | 19.6464 | 0.90002 | 21.0232 | 1.3768 | 0.065489 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 10 | Solution B | 224 | 2 | 8 | 10 | 44.8 | 1600.08 | 1777.6 | 81191.53 | 79591.45 | 49724.43 | 47946.83 | 1.66 | 42.3983 | 0.99858 | 43.5992 | 1.2008 | 0.027543 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 11 | Solution B | 224 | 3 | 7 | 10 | 67.2 | 1600.08 | 1777.6 | 126614.53 | 125014.45 | 76898.28 | 75120.68 | 1.6642 | 66.849 | 1.1431 | 67.0245 | 0.1755 | 0.0026184 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 12 | Solution B | 224 | 4 | 6 | 10 | 89.6 | 1600.08 | 1777.6 | 168237.15 | 166637.07 | 99941.93 | 98164.33 | 1.6975 | 91.7416 | -2.8926 | 90.6708 | 1.0708 | 0.01181 | Hakai | FM (no blank) - Inconsistent Values
2017-08-16 | Q5 | 720001154 | Quadra | 13 | Solution B | 224 | 5 | 5 | 10 | 112 | 1600.08 | 1777.6 | 207934.67 | 206334.59 | 124769.28 | 122991.68 | 1.6776 | 111.6651 | -0.34474 | 111.8325 | 0.16747 | 0.0014975 | Hakai | FM (no blank) - Inconsistent Values
2015-03-16 | C1 | 720000982 | Calvert | 1 | Solution C | 20.3 | 0.05 | 9.95 | 10 | 0.1015 | 803.06 |  | 736.5 | -66.56 | 673.41 | -129.65 | 0.51338 | 0.086155 | -0.22518 | 0.093828 | 0.0076724 | 0.081771 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 2 | Solution C | 20.3 | 0.1 | 9.9 | 10 | 0.203 | 803.06 |  | 1027.11 | 224.05 | 832.03 | 28.97 | 7.7339 | 0.2664 | -0.23533 | 0.2347 | 0.0317 | 0.13507 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 3 | Solution C | 20.3 | 0.25 | 9.75 | 10 | 0.5075 | 803.06 |  | 1395.51 | 592.45 | 1068.37 | 265.31 | 2.233 | 0.44674 | -0.16224 | 0.47712 | 0.03038 | 0.063674 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 4 | Solution C | 20.3 | 0.5 | 9.5 | 10 | 1.015 | 803.06 |  | 2224.97 | 1421.91 | 1526.62 | 723.56 | 1.9652 | 0.95366 | -0.17778 | 0.98433 | 0.03067 | 0.031158 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 5 | Solution C | 20.3 | 1 | 9 | 10 | 2.03 | 803.06 |  | 3878.41 | 3075.35 | 2357.2 | 1554.14 | 1.9788 | 2.0774 | -0.41082 | 2.0537 | 0.023676 | 0.011529 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 6 | Solution C | 20.3 | 2 | 8 | 10 | 4.06 | 803.06 |  | 7163.98 | 6360.92 | 4275.7 | 3472.64 | 1.8317 | 3.9442 | -0.22044 | 4.0021 | 0.057894 | 0.014466 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 7 | Solution C | 20.3 | 3 | 7 | 10 | 6.09 | 803.06 |  | 10116.15 | 9313.09 | 6185.66 | 5382.6 | 1.7302 | 5.3674 | 0.40441 | 5.7287 | 0.36128 | 0.063064 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 8 | Solution C | 20.3 | 4 | 6 | 10 | 8.12 | 803.06 |  | 13398.45 | 12595.39 | 8428.37 | 7625.31 | 1.6518 | 6.7871 | 1.3897 | 7.4536 | 0.66645 | 0.089414 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 9 | Solution B | 200 | 0.5 | 9.5 | 10 | 10 | 803.06 |  | 16346.27 | 15543.21 | 9360.24 | 8557.18 | 1.8164 | 9.5401 | -0.36405 | 9.77 | 0.22997 | 0.023538 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 10 | Solution B | 200 | 0.75 | 9.25 | 10 | 15 | 803.06 |  | 25867 | 25063.94 | 14512.24 | 13709.18 | 1.8283 | 15.506 | -0.80538 | 15.253 | 0.25298 | 0.016586 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 11 | Solution B | 200 | 1 | 9 | 10 | 20 | 803.06 |  | 34263.94 | 33460.88 | 19976.65 | 19173.59 | 1.7452 | 19.5106 | 1.0496 | 19.7553 | 0.2447 | 0.012386 | Hakai
2015-03-16 | C1 | 720000982 | Calvert | 12 | Solution B | 200 | 1.5 | 8.5 | 10 | 30 | 803.06 |  | 49826.06 | 49023 | 28662.02 | 27858.96 | 1.7597 | 28.9014 | 0.9722 | 29.4507 | 0.54928 | 0.018651 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 1 | Solution C | 22.8 | 0.1 | 9.9 | 10 | 0.228 | 718.67 | 711.98 | 854.17 | 135.5 | 744.72 | 32.74 | 4.1387 | 0.14302 | -0.107 | 0.18551 | 0.042489 | 0.22904 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 2 | Solution C | 22.8 | 0.25 | 9.75 | 10 | 0.57 | 718.67 | 711.98 | 1398.72 | 680.05 | 963.9 | 251.92 | 2.6995 | 0.59587 | -0.31874 | 0.58294 | 0.012936 | 0.022192 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 3 | Solution C | 22.8 | 0.5 | 9.5 | 10 | 1.14 | 718.67 | 711.98 | 2313.05 | 1594.38 | 1658.48 | 946.5 | 1.6845 | 0.90172 | 0.13952 | 1.0209 | 0.11914 | 0.1167 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 4 | Solution C | 22.8 | 1 | 9 | 10 | 2.28 | 718.67 | 711.98 | 4167.19 | 3448.52 | 2682.2 | 1970.22 | 1.7503 | 2.0575 | 0.10992 | 2.1688 | 0.11125 | 0.051296 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 5 | Solution C | 22.8 | 2 | 8 | 10 | 4.56 | 718.67 | 711.98 | 7695.88 | 6977.21 | 4552.82 | 3840.84 | 1.8166 | 4.3652 | -0.13993 | 4.4626 | 0.097395 | 0.021825 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 6 | Solution C | 22.8 | 3 | 7 | 10 | 6.84 | 718.67 | 711.98 | 11344.05 | 10625.38 | 6563.11 | 5851.13 | 1.816 | 6.6448 | -0.20804 | 6.7424 | 0.097591 | 0.014474 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 7 | Solution C | 22.8 | 4 | 6 | 10 | 9.12 | 718.67 | 711.98 | 14894.4 | 14175.73 | 8586.1 | 7874.12 | 1.8003 | 8.7706 | -0.10834 | 8.9453 | 0.1747 | 0.01953 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 8 | Solution C | 22.8 | 6 | 4 | 10 | 13.68 | 718.67 | 711.98 | 22703.64 | 21984.97 | 13316.25 | 12604.27 | 1.7442 | 13.0561 | 0.80977 | 13.368 | 0.31196 | 0.023336 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 9 | Solution B | 216.8 | 1 | 9 | 10 | 21.68 | 718.67 | 711.98 | 33616.33 | 32897.66 | 19188 | 18476.02 | 1.7806 | 20.0721 | 0.25324 | 20.876 | 0.80396 | 0.038511 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 10 | Solution B | 216.8 | 2 | 8 | 10 | 43.36 | 718.67 | 711.98 | 68864.16 | 68145.49 | 39017.64 | 38305.66 | 1.779 | 41.5312 | 0.60858 | 42.4456 | 0.91441 | 0.021543 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 11 | Solution B | 216.8 | 3 | 7 | 10 | 65.04 | 718.67 | 711.98 | 104264.17 | 103545.5 | 58454.19 | 57742.21 | 1.7932 | 63.7492 | -0.22742 | 64.3946 | 0.64541 | 0.010023 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 12 | Solution B | 216.8 | 4 | 6 | 10 | 86.72 | 718.67 | 711.98 | 141382.1 | 140663.43 | 78659.96 | 77947.98 | 1.8046 | 87.2876 | -1.5376 | 87.0038 | 0.28379 | 0.0032618 | Hakai
2016-04-16 | C2 | 720000982 | Calvert | 13 | Solution B | 216.8 | 5 | 5 | 10 | 108.4 | 718.67 | 711.98 | 177511.25 | 176792.58 | 97887.9 | 97175.92 | 1.8193 | 110.8107 | -3.9083 | 109.6054 | 1.2054 | 0.010997 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 1 | Solution C | 23.1 | 0.1 | 9.9 | 10 | 0.231 | 603.96 | 606.46 | 1038.74 | 434.78 | 868.42 | 261.96 | 1.6597 | 0.24859 | 0.039793 | 0.2398 | 0.0087951 | 0.036677 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 2 | Solution C | 23.1 | 0.25 | 9.75 | 10 | 0.5775 | 603.96 | 606.46 | 1332.77 | 728.81 | 1040.71 | 434.25 | 1.6783 | 0.42371 | 0.054346 | 0.5006 | 0.076897 | 0.15361 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 3 | Solution C | 23.1 | 0.5 | 9.5 | 10 | 1.155 | 603.96 | 606.46 | 2377.29 | 1773.33 | 1610.81 | 1004.35 | 1.7656 | 1.1061 | -0.00047129 | 1.1306 | 0.024437 | 0.021615 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 4 | Solution C | 23.1 | 1 | 9 | 10 | 2.31 | 603.96 | 606.46 | 4007.57 | 3403.61 | 2534.43 | 1927.97 | 1.7654 | 2.1226 | -0.00017236 | 2.2163 | 0.093695 | 0.042275 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 5 | Solution C | 23.1 | 2 | 8 | 10 | 4.62 | 603.96 | 606.46 | 7569.27 | 6965.31 | 4501.95 | 3895.49 | 1.788 | 4.4157 | -0.12732 | 4.5179 | 0.10213 | 0.022607 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 6 | Solution C | 23.1 | 3 | 7 | 10 | 6.93 | 603.96 | 606.46 | 11350.08 | 10746.12 | 6750.96 | 6144.5 | 1.7489 | 6.6191 | 0.14515 | 6.7746 | 0.15544 | 0.022944 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 7 | Solution C | 23.1 | 4 | 6 | 10 | 9.24 | 603.96 | 606.46 | 14337.37 | 13733.41 | 8366.23 | 7759.77 | 1.7698 | 8.5927 | -0.050213 | 8.9163 | 0.32366 | 0.036299 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 8 | Solution C | 23.1 | 3 | 2 | 5 | 13.86 | 603.96 | 606.46 | 22797.86 | 22193.9 | 13353.62 | 12747.16 | 1.7411 | 13.5885 | 0.44441 | 13.7243 | 0.13575 | 0.0098909 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 9 | Solution B | 226.6 | 1 | 9 | 10 | 22.66 | 603.96 | 606.46 | 37085.22 | 36481.26 | 21161.97 | 20555.51 | 1.7748 | 22.9081 | -0.27926 | 22.7841 | 0.12407 | 0.0054454 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 10 | Solution B | 226.6 | 2 | 8 | 10 | 45.32 | 603.96 | 606.46 | 70505.28 | 69901.32 | 40291.04 | 39684.58 | 1.7614 | 43.4648 | 0.22265 | 44.3924 | 0.92761 | 0.020896 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 11 | Solution B | 226.6 | 3 | 7 | 10 | 67.98 | 603.96 | 606.46 | 112496.93 | 111892.97 | 63418.94 | 62812.48 | 1.7814 | 70.599 | -1.4509 | 69.2895 | 1.3095 | 0.018899 | Hakai
2017-05-02 | C3 | 720000982 | Calvert | 12 | Solution B | 226.6 | 4 | 6 | 10 | 90.64 | 603.96 | 606.46 | 144014.57 | 143410.61 | 82239.46 | 81633 | 1.7568 | 88.863 | 1.0041 | 89.7515 | 0.8885 | 0.0098996 | Hakai
2018-04-17 | Q6 | 720001154 | Quadra | 1 | C | 19.5 | 0.5 | 9.5 | 10 | 0.975 | 1343.78 | 1343.65 | 3545.66 | 2201.88 | 2470.25 | 1126.6 | 1.9544 | 1.2227 | -0.35737 | 1.0989 | 0.12385 | 0.11271 | Hakai | High FM
2018-04-17 | Q6 | 720001154 | Quadra | 2 | C | 19.5 | 1 | 9 | 10 | 1.95 | 1343.78 | 1343.65 | 5496.01 | 4152.23 | 3556.11 | 2212.46 | 1.8767 | 2.2057 | -0.50635 | 2.0779 | 0.12786 | 0.061535 | Hakai | High FM
2018-04-17 | Q6 | 720001154 | Quadra | 3 | C | 19.5 | 2 | 8 | 10 | 3.9 | 1343.78 | 1343.65 | 9437.68 | 8093.9 | 6320.57 | 4976.92 | 1.6263 | 3.5443 | 0.2784 | 3.7222 | 0.17783 | 0.047777 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 4 | C | 19.5 | 3 | 7 | 10 | 5.85 | 1343.78 | 1343.65 | 13943.41 | 12599.63 | 8786.71 | 7443.06 | 1.6928 | 5.8636 | -0.1466 | 5.8568 | 0.0067795 | 0.0011575 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 5 | C | 19.5 | 4 | 6 | 10 | 7.8 | 1343.78 | 1343.65 | 18021.27 | 16677.49 | 11336.88 | 9993.23 | 1.6689 | 7.6007 | 0.075017 | 7.7004 | 0.099649 | 0.012941 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 6 | C | 19.5 | 6 | 4 | 10 | 11.7 | 1343.78 | 1343.65 | 26079.12 | 24735.34 | 16218.28 | 14874.63 | 1.6629 | 11.2127 | 0.21243 | 11.4563 | 0.24367 | 0.02127 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 7 | B | 194 | 1 | 9 | 10 | 19.4 | 1343.78 | 1343.65 | 40681.05 | 39337.27 | 24817.96 | 23474.31 | 1.6758 | 18.0378 | -0.0074146 | 18.7189 | 0.68108 | 0.036384 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 8 | B | 194 | 2 | 8 | 10 | 38.8 | 1343.78 | 1343.65 | 84581.1 | 83237.32 | 50765.36 | 49421.71 | 1.6842 | 38.4519 | -0.49146 | 38.6259 | 0.17406 | 0.0045063 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 9 | B | 194 | 3 | 7 | 10 | 58.2 | 1343.78 | 1343.65 | 126441.43 | 125097.65 | 75298.01 | 73954.36 | 1.6916 | 58.1553 | -1.3515 | 58.1776 | 0.022367 | 0.00038446 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720001154 | Quadra | 10 | B | 194 | 4 | 6 | 10 | 77.6 | 1343.78 | 1343.65 | 171518.35 | 170174.57 | 101362.88 | 100019.23 | 1.7014 | 79.774 | -2.95 | 78.687 | 1.087 | 0.013814 | Hakai | FM below Threshold (1.7)
2018-04-17 | Q6 | 720000982 | Calvert | 1 | C | 19.5 | 0.5 | 9.5 | 10 | 0.975 | 436.77 | 461.39 | 2099.23 | 1662.46 | 1372.51 | 911.12 | 1.8246 | 0.98249 | -0.035984 | 0.97874 | 0.003744 | 0.0038253 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 2 | C | 19.5 | 1 | 9 | 10 | 1.95 | 436.77 | 461.39 | 3742.46 | 3305.69 | 2277.32 | 1815.93 | 1.8204 | 1.9481 | -0.061629 | 1.949 | 0.00095945 | 0.00049227 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 3 | C | 19.5 | 2 | 8 | 10 | 3.9 | 436.77 | 461.39 | 7138.75 | 6701.98 | 4183.65 | 3722.26 | 1.8005 | 3.8964 | -0.029608 | 3.8982 | 0.0017881 | 0.00045871 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 4 | C | 19.5 | 3 | 7 | 10 | 5.85 | 436.77 | 461.39 | 10556.59 | 10119.82 | 6116.01 | 5654.62 | 1.7897 | 5.8389 | 0.035312 | 5.8445 | 0.0055459 | 0.00094892 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 5 | C | 19.5 | 4 | 6 | 10 | 7.8 | 436.77 | 461.39 | 13835.89 | 13399.12 | 8008.42 | 7547.03 | 1.7754 | 7.6525 | 0.18765 | 7.7262 | 0.073764 | 0.0095472 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 6 | C | 19.5 | 6 | 4 | 10 | 11.7 | 436.77 | 461.39 | 20522.39 | 20085.62 | 11696.99 | 11235.6 | 1.7877 | 11.5727 | 0.099232 | 11.6364 | 0.063646 | 0.0054696 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 7 | B | 194 | 1 | 9 | 10 | 19.4 | 436.77 | 461.39 | 31993.25 | 31556.48 | 18231.71 | 17770.32 | 1.7758 | 18.0274 | 0.433 | 18.7137 | 0.68628 | 0.036673 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 8 | B | 194 | 2 | 8 | 10 | 38.8 | 436.77 | 461.39 | 67107.86 | 66671.09 | 37650.29 | 37188.9 | 1.7928 | 38.5523 | 0.080832 | 38.6762 | 0.12384 | 0.003202 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 9 | B | 194 | 3 | 7 | 10 | 58.2 | 436.77 | 461.39 | 100198.93 | 99762.16 | 56258 | 55796.61 | 1.788 | 57.4914 | 0.47205 | 57.8457 | 0.35428 | 0.0061245 | Hakai
2018-04-17 | Q6 | 720000982 | Calvert | 10 | B | 194 | 4 | 6 | 10 | 77.6 | 436.77 | 461.39 | 134841.78 | 134405.01 | 75568.9 | 75107.51 | 1.7895 | 77.5402 | 0.48411 | 77.5701 | 0.029883 | 0.00038524 | Hakai
2018-05-04 | Q6 | 720001154 | Quadra | 1 | C | 19.5 | 0.5 | 9.5 | 10 | 0.9979 | 1124.29 | 1114.55 | 3877.29 | 2753 | 2806.53 | 1691.98 | 1.6271 | 1.2629 | 0.20843 | 1.1304 | 0.13252 | 0.11723 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 2 | C | 19.5 | 1 | 9 | 10 | 7.9832 | 1124.29 | 1114.55 | 16855.66 | 15731.37 | 10243.43 | 9128.88 | 1.7233 | 7.859 | 0.079597 | 7.9211 | 0.062118 | 0.0078421 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 3 | C | 19.5 | 2 | 8 | 10 | 25.4473 | 1124.29 | 1114.55 | 53696.7 | 52572.41 | 31649.18 | 30534.63 | 1.7217 | 26.2316 | 0.32157 | 25.8395 | 0.39217 | 0.015177 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 4 | C | 19.5 | 3 | 7 | 10 | 49.896 | 1124.29 | 1114.55 | 97209.19 | 96084.9 | 57275.46 | 56160.91 | 1.7109 | 47.5217 | 1.3164 | 48.7088 | 1.1872 | 0.024373 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 5 | C | 19.5 | 4 | 6 | 10 | 74.844 | 1124.29 | 1114.55 | 146783.65 | 145659.36 | 85821.94 | 84707.39 | 1.7196 | 72.5513 | 1.111 | 73.6977 | 1.1463 | 0.015555 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 6 | C | 19.5 | 6 | 4 | 10 | 99.792 | 1124.29 | 1114.55 | 199767.51 | 198643.22 | 115294.43 | 114179.88 | 1.7397 | 100.537 | -1.2451 | 100.1645 | 0.37249 | 0.0037188 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 7 | B | 194 | 1 | 9 | 10 | 207.9 | 1124.29 | 1114.55 | 421677.96 | 420553.67 | 241713.62 | 240599.07 | 1.7479 | 214.2005 | -4.9733 | 211.0503 | 3.1503 | 0.014927 | DFO
2018-05-04 | Q6 | 720001154 | Quadra | 8 | B | 194 | 2 | 8 | 10 | 299.376 | 1124.29 | 1114.55 | 592881 | 591756.71 | 339080.4 | 337965.85 | 1.7509 | 302.0881 | -8.1897 | 300.732 | 1.356 | 0.0045091 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 1 | C | 19.5 | 0.5 | 9.5 | 10 | 1.23 | 393.38 | 386.46 | 2803.42 | 2410.04 | 1784.74 | 1398.28 | 1.7236 | 1.2074 | 0.028193 | 1.2187 | 0.011298 | 0.0092707 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 2 | C | 19.5 | 1 | 9 | 10 | 7.67 | 393.38 | 386.46 | 14812.75 | 14419.37 | 8832.83 | 8446.37 | 1.7072 | 7.128 | 0.33568 | 7.399 | 0.271 | 0.036627 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 3 | C | 19.5 | 2 | 8 | 10 | 26.69 | 393.38 | 386.46 | 54494.58 | 54101.2 | 31415.45 | 31028.99 | 1.7436 | 27.5337 | -0.11476 | 27.1118 | 0.42184 | 0.015559 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 4 | C | 19.5 | 3 | 7 | 10 | 53.38 | 393.38 | 386.46 | 100196.39 | 99803.01 | 57952.61 | 57566.15 | 1.7337 | 50.4042 | 0.46442 | 51.8921 | 1.4879 | 0.028673 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 5 | C | 19.5 | 4 | 6 | 10 | 83.4 | 393.38 | 386.46 | 162408.43 | 162015.05 | 93587.53 | 93201.07 | 1.7383 | 82.1205 | 0.23706 | 82.7603 | 0.63974 | 0.0077301 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 6 | C | 19.5 | 6 | 4 | 10 | 240.19 | 393.38 | 386.46 | 466241.78 | 465848.4 | 263958.68 | 263572.22 | 1.7674 | 241.3902 | -8.4834 | 240.7901 | 0.60012 | 0.0024923 | DFO
2019-05-09 | Q6 | 720001154 | Quadra | 7 | B | 194 | 1 | 9 | 10 | 306.91 | 393.38 | 386.46 | 608900.62 | 608507.24 | 344274.87 | 343888.41 | 1.7695 | 315.7881 | -11.9094 | 311.349 | 4.439 | 0.014257 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 1 | C | 19.5 | 0.5 | 9.5 | 10 | 0.9979 | 439.68 | 446.11 | 2380.73 | 1941.05 | 1540.2 | 1094.09 | 1.7741 | 1.1973 | 0.088453 | 1.0976 | 0.09969 | 0.090826 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 2 | C | 19.5 | 1 | 9 | 10 | 7.9832 | 439.68 | 446.11 | 12643.25 | 12203.57 | 7121.85 | 6675.74 | 1.828 | 7.8143 | 0.030821 | 7.8987 | 0.084475 | 0.010695 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 3 | C | 19.5 | 2 | 8 | 10 | 25.4473 | 439.68 | 446.11 | 40975.03 | 40535.35 | 22771.33 | 22325.22 | 1.8157 | 25.7422 | 0.49353 | 25.5948 | 0.14745 | 0.005761 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 4 | C | 19.5 | 3 | 7 | 10 | 49.896 | 439.68 | 446.11 | 76034.14 | 75594.46 | 42055.57 | 41609.46 | 1.8168 | 48.0419 | 0.85595 | 48.9689 | 0.92706 | 0.018932 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 5 | C | 19.5 | 4 | 6 | 10 | 74.844 | 439.68 | 446.11 | 115605.78 | 115166.1 | 62879.03 | 62432.92 | 1.8446 | 74.5447 | -1.1759 | 74.6943 | 0.14966 | 0.0020037 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 6 | C | 19.5 | 6 | 4 | 10 | 99.792 | 439.68 | 446.11 | 154432.07 | 153992.39 | 83019.63 | 82573.52 | 1.8649 | 100.9591 | -3.9219 | 100.3756 | 0.58357 | 0.0058138 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 7 | B | 194 | 1 | 9 | 10 | 207.9 | 439.68 | 446.11 | 326116.06 | 325676.38 | 176566.34 | 176120.23 | 1.8492 | 211.4156 | -4.4459 | 209.6578 | 1.7578 | 0.008384 | DFO
2018-05-04 | C4 | 720000982 | Calvert | 8 | B | 194 | 2 | 8 | 10 | 299.376 | 439.68 | 446.11 | 465952.21 | 465512.53 | 251101.68 | 250655.57 | 1.8572 | 303.7261 | -9.1654 | 301.551 | 2.175 | 0.0072128 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 1 | - | - | - | - | 10 | 1.23 | 371.85 | 389.03 | 2195.03 | 1823.18 | 1384.61 | 995.58 | 1.8313 | 1.1933 | 0.022406 | 1.2116 | 0.018372 | 0.015163 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 2 | - | - | - | - | 10 | 7.67 | 371.85 | 389.03 | 11913.05 | 11541.2 | 6654.31 | 6265.28 | 1.8421 | 7.607 | 0.043313 | 7.6385 | 0.03152 | 0.0041265 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 3 | - | - | - | - | 10 | 26.69 | 371.85 | 389.03 | 41966.06 | 41594.21 | 23185.69 | 22796.66 | 1.8246 | 27.1028 | 0.73325 | 26.8964 | 0.2064 | 0.0076739 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 4 | - | - | - | - | 10 | 53.38 | 371.85 | 389.03 | 78856.76 | 78484.91 | 43047.32 | 42658.29 | 1.8399 | 51.6558 | 0.43251 | 52.5179 | 0.86212 | 0.016416 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 5 | - | - | - | - | 10 | 83.4 | 371.85 | 389.03 | 126385.75 | 126013.9 | 68473.46 | 68084.43 | 1.8508 | 83.5242 | -0.38916 | 83.4621 | 0.062117 | 0.00074425 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 6 | - | - | - | - | 10 | 240.19 | 371.85 | 389.03 | 363018 | 362646.15 | 194131.06 | 193742.03 | 1.8718 | 243.5304 | -6.96 | 241.8602 | 1.6702 | 0.0069057 | DFO
2019-05-09 | C4 | 720000982 | Calvert | 7 | - | - | - | - | 10 | 306.91 | 371.85 | 389.03 | 464944.34 | 464572.49 | 249123.14 | 248734.11 | 1.8677 | 311.2015 | -7.4825 | 309.0557 | 2.1457 | 0.0069428 | DFO